Hardware Parameters and Reproducibility Support

2D LiDAR to 3D Point Cloud Reconstruction Framework

# Purpose

This document summarizes the hardware parameters required to reproduce the proposed 2D LiDAR-based 3D point cloud reconstruction framework. It is intended to be included in the public repository under docs/hardware_parameters.md or docs/hardware_parameters.pdf/docx, and referenced in the manuscript Code and Data Availability section.

## Table 1. System-level configuration

| Parameter | Value / Setting | Role in reconstruction |
| --- | --- | --- |
| Acquisition principle | Sequential 2D LiDAR scans stacked along the vertical axis | Generates 3D point clouds by assigning a vertical coordinate to each completed 360-degree scan layer. |
| Scanning geometry | 360-degree horizontal scan with controlled vertical translation | Preserves the native 2D LiDAR scan plane while extending acquisition into 3D. |
| Primary output | Raw and filtered 3D point cloud coordinates (X, Y, Z) | Used for visualization, filtering, and quantitative evaluation. |
| Operating scene assumption | Mostly static scene during acquisition | Required because the complete 3D point cloud is accumulated over multiple scan layers. |
| Reference frame | LiDAR-centered frame with vertical Z-axis | X and Y are computed from range and angle; Z is obtained from vertical displacement. |

## Table 2. LiDAR sensor parameters

| Parameter | Value / Setting | Notes |
| --- | --- | --- |
| Sensor model | SLAMTEC RPLIDAR A1 | Low-cost 2D spinning LiDAR used for planar range acquisition. |
| Measurement principle | Laser triangulation | Produces distance and angle measurements for each scan point. |
| Angular coverage | 0-360 degrees | Full planar scan used for each reconstructed layer. |
| Typical scan frequency | 1-10 Hz; representative value around 5.5 Hz | Used for synchronization between scan completion and vertical motion. |
| Measurement rate | More than 8000 measurements/s | Reported sensor capability. |
| Approximate points per revolution | Around 1450 points/revolution | Depends on motor speed and acquisition mode. |
| Distance unit | Millimeters | Used in coordinate reconstruction and error evaluation. |
| Angular unit | Degrees | Converted to radians for trigonometric coordinate transformation. |
| Minimum measurable distance | Approximately 150 mm | Points below this range should be filtered. |
| Distance resolution | Below 0.5 mm for targets closer than 1.5 m | Manufacturer-reported behavior; actual accuracy depends on surface and environment. |
| Relative ranging error | Below 1% over operating range | Reported sensor behavior. |
| Laser wavelength | Approximately 785 nm | Infrared laser source. |
| Laser power | Typically around 3 mW; up to 5 mW | Reported low-power operation. |
| Laser safety class | Class 1 | Eye-safe classification. |
| Supply voltage | 5 V | Stable power supply is required. |
| Current consumption | Approximately 80 mA standby; 350 mA normal operation | Startup current may be higher. |
| Environmental condition | Primarily indoor; avoid direct sunlight | Strong ambient light may degrade measurement quality. |

## Table 3. Vertical translation and motion parameters

| Parameter | Value / Setting | Notes |
| --- | --- | --- |
| Vertical motion type | Controlled linear translation along Z-axis | Used to assign scan layers to vertical positions. |
| Guiding mechanism | Linear sliding rails | Reduces lateral play and improves layer consistency. |
| Transmission mechanism | Trapezoidal lead screw with nut carriage | Converts motor rotation into vertical displacement. |
| Lead screw start type | Single-start screw | Pitch equals linear advance per full revolution. |
| Lead screw pitch | TBD mm/revolution | Must be filled using the exact prototype screw specification. |
| Usable vertical stroke | TBD mm | Must be filled from measured prototype travel range. |
| Vertical step size | TBD mm/layer | Defines spacing between successive 2D scan layers. |
| Number of vertical layers | TBD | Depends on stroke and selected vertical step size. |
| Vertical speed | TBD mm/s | Required for synchronization-error analysis. |
| Backlash estimate | TBD mm | Should be measured or bounded for error propagation analysis. |

## Table 4. Motor, driver, and embedded-control parameters

| Parameter | Value / Setting | Notes |
| --- | --- | --- |
| Motor type | NEMA 17 stepper motor | Selected for repeatable discrete positioning and holding capability. |
| Motor wiring | Bipolar 4-wire | Compatible with TB6600-type driver. |
| Nominal step angle | TBD degrees/step | Common value is often 1.8 degrees, but the exact motor specification must be confirmed. |
| Steps per revolution | TBD steps/revolution | Derived from the step angle and microstepping setting. |
| Driver | TB6600 stepper motor driver | Provides adjustable current and microstep settings. |
| Drive mode | Full-step emphasized in manuscript; microstepping setting TBD | Full-step prioritizes torque and holding stability. |
| Microcontroller | Arduino UNO | Generates step and direction signals. |
| Control signals | STEP and DIR | One pulse advances the motor by the configured step increment. |
| Human interface | Push-button start/stop | Used to initiate or interrupt acquisition sequences. |
| Timing control | Microsecond-level pulse delays | Controls motor speed through pulse frequency. |

## Table 5. Acquisition and synchronization parameters

| Parameter | Value / Setting | Notes |
| --- | --- | --- |
| Scan boundary detection | Start-of-scan synchronization marker in LiDAR data stream | Used to identify complete 360-degree rotations. |
| Rotation-time estimation | Time between consecutive start-of-scan markers | Used to monitor LiDAR speed and temporal consistency. |
| Layer assignment | One completed scan associated with one vertical position | Ensures each horizontal slice has a unique Z coordinate. |
| Incomplete scans | Discarded or not associated with a layer | Prevents reconstruction artifacts. |
| Synchronization delay | TBD s | Needed for error propagation using delta_z_sync = v_z * delta_t. |
| Host-side processing | LiDAR acquisition, scan ordering, coordinate conversion, and file writing | Performed on the computer using the RPLIDAR SDK and reconstruction code. |

## Table 6. Coordinate reconstruction parameters

| Parameter | Formula / Setting | Notes |
| --- | --- | --- |
| Input LiDAR measurement | Range r and angle theta | Each 2D point is represented in polar coordinates. |
| X coordinate | x = r cos(theta) | theta must be converted to radians in implementation. |
| Y coordinate | y = r sin(theta) | Computed in the LiDAR scan plane. |
| Z coordinate | z = vertical_position | Assigned from the controlled vertical displacement. |
| Output point | P = (x, y, z) | Stored as reconstructed 3D point cloud coordinate. |
| Invalid point filtering | Distance below minimum range or low-quality returns removed | Improves point cloud reliability. |

## Table 7. Recommended repository files for hardware reproducibility

| File | Recommended location | Purpose |
| --- | --- | --- |
| hardware_parameters.md | docs/hardware_parameters.md | Human-readable hardware parameter summary. |
| hardware_parameters.csv | docs/hardware_parameters.csv | Machine-readable table for reproducibility. |
| wiring_diagram.png | docs/figures/wiring_diagram.png | Arduino, driver, motor, and LiDAR connection diagram. |
| calibration_config.yaml | config/calibration_config.yaml | Stores calibrated vertical step size, offsets, and correction parameters. |
| acquisition_config.yaml | config/acquisition_config.yaml | Stores serial ports, baud rates, scan mode, and synchronization settings. |
| sample_scan.xyz or .csv | datasets/sample_scans/ | Representative scan used to verify the reconstruction pipeline. |